# Release Notes  —  v1.0

2D/3D Fluoroscopy Registration Pipeline
October 2026

✓  4.3 mm PDE on real cadaveric fluoroscopy (DeepFluoro)  |  ✓  0.32 mm PDE on public spine CT (VerSe19)  |  ✓  Zero manual C-arm intervention  |  ✓  Zero hardware cost

# 1.  What Is This Release?

This is the first public release of the NIMS 2D/3D fluoroscopy registration pipeline, made available for evaluation by Neurosapphire / R-Zoom Technologies. It delivers a complete, validated registration system ready for research and integration evaluation.

# 2.  Validated Results

| Test | PDE | Dataset | C-arm input | Status |
| --- | --- | --- | --- | --- |
| Pelvis, 2-view, real fluoroscopy | 4.3 mm | DeepFluoro (Johns Hopkins cadaveric) | None | ✓ Validated |
| Spine, 2-view, synthetic DRR | 0.32 mm | VerSe19 (public lumbar CT) | None | ✓ Validated |
| Spine, 2-view, real C-arm (NIMS) | Under evaluation | NIMS surgical dataset | None (blind) | ⚠ Gaps identified |

# 3.  What Is Included in This Release

## pipeline/  —  Ready-to-Run Scripts

sanity_check_render.py — verify DRR renderer on your hardware

build_drr_bank_deepfluoro.py — build DeepFluoro DRR bank

run_deepfluoro_multiview_test.py — reproduce 4.3 mm result

build_drr_bank.py — generic bank builder for any CT

run_verse_multiview_test.py — reproduce 0.32 mm result

run_nims_multiview_test.py — NIMS real dataset test (update paths)

run_nims_auto_pedicle_test.py — NIMS test with auto pedicle extraction

## src/  —  Core Library

deepfluoro_drr.py — GPU Beer-Lambert DRR renderer

similarity.py — Gradient Orientation similarity metric

metal_suppression.py — surgical instrument inpainting

camera_config.py — C-arm geometry configuration loader

landmark_model.py — LandmarkHeatmapNet (MobileNetV3-Small + UNet, 18-class)

landmark_detector.py — inference wrapper for landmark CNN

cyclegan.py — CycleGAN for X-ray → DRR style transfer

optimizer.py — CMA-ES multi-start optimiser

generic_patient_loader.py — patient data folder loader

## scripts/  —  Training and Data Preparation

train_landmark_detector.py — LOPO CNN training

train_cyclegan.py — CycleGAN training

generate_synthetic_drrs.py — synthetic DRR generation from CT

extract_transverse_process_tips.py — CT pedicle auto-extraction

convert_xrays_to_drr.py — CycleGAN inference on real X-rays

# 4.  Fixes Applied in This Release

| Item | Description |
| --- | --- |
| Camera config corrected | ziehm_vision_fd.json: SID 1110→1050mm, pixel_spacing 0.2→0.288mm/px. Source: verified directly from DICOM tags. |
| Metal suppression improved | Threshold p95/0.75 → p97/0.82, dilation 13px → 21px, removed early-exit skip that prevented suppression of thin K-wires. |
| Training crash fixed | train_landmark_detector.py: added redirect_stderr to suppress timm backbone warning that was aborting PowerShell training runs at 0 epochs. |
| Patient names removed | All patient identifiers removed from source code and replaced with generic identifiers (PATIENT_001, NIMS dataset) for privacy and portability. |
| Hardcoded paths replaced | NIMS pipeline scripts now have a clear USER CONFIGURATION section at the top for setting data paths, instead of buried hardcoded absolute paths. |
| Manuscript removed | LaTeX manuscript files removed from the public release build. |
| Unused code removed | Old SSM regressor, pose regressor, deformable registration, and all internal dev/test scripts removed from the release. |

# 5.  Known Gaps and What Is Needed to Fix Them

| Gap | Root Cause | Fix Required | Cost |
| --- | --- | --- | --- |
| Spine landmark collinearity | 3 AP centroids (L3-L5) near-collinear in 3D (PCA ratio 0.128) | Option A: Radiopaque skin sticker markers placed before CT Option B: Retrain CNN with pedicle detections | Option A: ~Rs 500/patient Option B: 13 patient datasets |
| C-arm isocenter calibration | Isocenter 3D position not independently measured; only angles from DICOM | One-time phantom calibration session with C-arm engineer (30 min, done once per machine) | Engineer time only |
| CNN training data | ~30 real X-ray images vs 7,200 synthetic DRRs; domain mismatch | Collect 13+ patient datasets with 2D labels; Retrain with CycleGAN rows | 2 min annotation per image; zero hardware cost |

# 6.  Not Included in This Release

Trained model checkpoints (.pth files) — too large for Git. Contact NIMS team.

Patient CT/X-ray data — not shareable. Use public datasets for testing.

DICOM integration layer — planned for v1.1.

Real-time surgeon display overlay — planned for v1.2.

Neurosapphire-specific integration layer — to be developed with R-Zoom team.

# 7.  Dependencies

| Package | Version | Purpose |
| --- | --- | --- |
| torch | 2.4.1 | Deep learning framework (GPU CUDA required for DRR speed) |
| torchvision | 0.19.1 | Image transforms + pretrained backbone weights |
| timm | 1.0.26 | MobileNetV3-Small backbone for LandmarkHeatmapNet |
| SimpleITK / pynrrd | 2.5.6 / 1.1.3 | Medical imaging I/O (CT volumes) |
| scipy | 1.10.1 | Nelder-Mead optimiser + signal processing |
| opencv-python-headless | 4.13.0.92 | Image processing, CLAHE, inpainting |
| numpy | 1.24.4 | Array operations |
| pydicom | latest | Reading C-arm DICOM frames |