# 2D/3D Fluoroscopy Registration Pipeline

Technical Documentation  |  Release v1.0  |  October 2026

# 1.  System Overview

This pipeline automatically registers a patient's pre-operative CT scan to an intraoperative C-arm fluoroscopy X-ray. It estimates the 3D pose of the spine or pelvis anatomy relative to the C-arm camera — enabling real-time surgical navigation without any manual annotation on the C-arm images and without any additional hardware.

Validated results:

| Test | Dataset | PDE Result | Target | Manual C-arm input |
| --- | --- | --- | --- | --- |
| Pelvis, real fluoroscopy | DeepFluoro (Johns Hopkins, cadaveric) | 4.3 mm | < 5 mm | None |
| Spine, synthetic DRR | VerSe19 (public lumbar CT) | 0.32 mm | < 2 mm | None |

# 2.  Repository Structure

| File / Folder | Purpose |
| --- | --- |
| README.md | Quick-start guide and overview |
| RELEASE_NOTES.md | What is in this release and known gaps |
| requirements.txt | Python package dependencies |
| configs/cameras/*.json | C-arm geometry configurations (SID, pixel spacing) |
| pipeline/ | Runnable end-to-end registration scripts |
| src/ | Core library modules (DRR renderer, GO metric, etc.) |
| scripts/ | Training and data-preparation scripts |
| docs/ | Technical documentation and release notes |
| data/sample/README_DATA.md | Instructions for placing datasets |

# 3.  Core Library Modules  (src/)

## 3.1  drr_renderer.py  /  deepfluoro_drr.py  —  DRR Renderer

Renders Digitally Reconstructed Radiographs (DRRs) from a CT volume at any camera pose using GPU-accelerated Beer-Lambert line integration. Takes a CTVolume object and a camera (source position, principal axis, u/v axes, SID) and produces a grayscale synthetic X-ray image.

Key classes and functions:

CTVolume(array, spacing, origin, hu_threshold) — wraps CT array, clips HU, prepares for rendering.

DRRRenderer(ct_volume) — stateful renderer; call .render(src, pax, u, v, size, pix_mm, sdd_mm).

render() returns a float32 [0,1] image. GPU required for acceptable speed (> 5 fps).

## 3.2  similarity.py  —  GO Similarity Metric

Implements Gradient Orientation (GO) similarity (Ketcha 2017). Measures how well a synthetic DRR matches a real X-ray by comparing gradient orientations at high-gradient pixels. Returns a score in [0, 1] — higher is better match.

Key functions:

gradient_orientation_similarity(drr, xray, sigma=2.0) — main GO metric.

go_cost(drr, xray) — returns 1 - similarity; used as minimisation cost.

fast_combined_cost(drr, xray) — blends GO (0.6) + edge-NCC (0.4) for robustness.

## 3.3  metal_suppression.py  —  Surgical Instrument Handling

Detects and inpaints bright surgical instruments (K-wires, catheters, retractors) that would otherwise dominate GO similarity. Two modes:

suppress_metal(xray) — blind adaptive threshold (p97 floor 0.82) + Telea inpainting.

suppress_metal_drr(xray, drr) — DRR-guided: masks pixels much brighter than DRR prediction.

## 3.4  camera_config.py  —  C-arm Camera Geometry

Loads camera geometry from JSON files in configs/cameras/. Returns a CameraConfig namedtuple with SID (mm), pixel_spacing_mm, image_size_px, and derived focal length (Fx, Fy).

load_camera_config(name_or_path) — loads by machine name or JSON file path.

Three built-in configs: ziehm_vision_fd, siemens_cios_spin, ge_oec_9900.

Add new machines by creating a JSON file in configs/cameras/.

## 3.5  landmark_model.py  —  LandmarkHeatmapNet

MobileNetV3-Small backbone + 4-stage UNet decoder. Detects 18 landmark classes (6 vertebra centroids D12/L1-L5 + 12 bilateral pedicle tips) in a 512x512 X-ray. Output: (B, 18, 128, 128) heatmap logits.

LandmarkHeatmapNet(pretrained=True, dropout=0.3) — build model.

HeatmapLoss(pos_weight=100.0) — weighted MSE on sigmoid heatmaps.

Total ~2.38M trainable parameters. Fits in 4 GB GPU.

## 3.6  landmark_detector.py  —  Inference Wrapper

Wraps LandmarkHeatmapNet for inference. Loads a checkpoint, runs the model on a preprocessed X-ray, extracts peak coordinates per class above a confidence threshold.

LandmarkDetector(ckpt_path, device) — load checkpoint and prepare for inference.

predict(xray_float32) — returns dict {name: [u, v]} for all detected landmarks.

## 3.7  cyclegan.py  —  Domain Style Transfer

CycleGAN (Zhu et al. 2017) for unpaired X-ray ↔ synthetic DRR style transfer. UNet generator (ngf=32, 4 ResBlocks) + 70×70 PatchGAN discriminator (ndf=32). Used to convert real X-rays to DRR style for reducing domain gap during landmark detection.

CycleGAN.load(path, device) — load trained checkpoint.

xray_to_drr(tensor) — apply G_AB generator (X-ray → DRR style).

## 3.8  optimizer.py  —  CMA-ES Multi-start Optimizer

Implements the Ketcha 2017 4-stage CMA-ES registration framework. In practice the pipeline uses scipy Nelder-Mead directly with multi-start (10 starts from top-10 GO retrieval candidates). optimizer.py provides the full CMA-ES alternative for future use.

# 4.  Pipeline Scripts  (pipeline/)

| Script | Purpose | Run command |
| --- | --- | --- |
| sanity_check_render.py | Quick test that DRR renderer works on your GPU. Renders a DRR and saves it. | python pipeline/sanity_check_render.py |
| build_drr_bank_deepfluoro.py | Builds a 500-entry DRR bank from the DeepFluoro CT at calibrated camera poses. | python pipeline/build_drr_bank_deepfluoro.py |
| run_deepfluoro_multiview_test.py | Full 2-view registration test on DeepFluoro. Produces the 4.3mm result. | python pipeline/run_deepfluoro_multiview_test.py |
| build_drr_bank.py | Generic DRR bank builder for any CT + camera config. Configurable via argparse. | python pipeline/build_drr_bank.py --help |
| run_verse_multiview_test.py | Full 2-view spine registration on VerSe19. Produces the 0.32mm result. | python pipeline/run_verse_multiview_test.py |
| run_nims_multiview_test.py | NIMS real patient 2-view test. Requires labeled AP+LAT sets for calibration. | Update paths at top, then: python pipeline/run_nims_multiview_test.py |
| run_nims_auto_pedicle_test.py | NIMS test with auto-extracted CT pedicle landmarks for improved geometry. | Update paths at top, then: python pipeline/run_nims_auto_pedicle_test.py |

# 5.  Training Scripts  (scripts/)

| Script | Purpose | Notes |
| --- | --- | --- |
| train_landmark_detector.py | LOPO cross-validation training of LandmarkHeatmapNet (18-class heatmap CNN). | Requires: data/landmark_manifest_real.csv + synthetic DRR manifest. Run: .\run_pedicle_training.ps1 or python scripts/train_landmark_detector.py --help |
| train_cyclegan.py | CycleGAN training for X-ray→DRR style transfer. | Requires: real X-ray images + synthetic DRRs. Run: python scripts/train_cyclegan.py --help |
| generate_synthetic_drrs.py | Renders 500-1000 synthetic DRRs per patient from CT. GPU required. | Run: python scripts/generate_synthetic_drrs.py --help |
| extract_transverse_process_tips.py | Auto-extracts pedicle tip positions from CT using HU thresholding. | Run: python scripts/extract_transverse_process_tips.py --help |
| convert_xrays_to_drr.py | Applies trained CycleGAN to convert real X-ray images to DRR style. | Requires trained CycleGAN checkpoint. Run: python scripts/convert_xrays_to_drr.py --help |

# 6.  Camera Configuration Files  (configs/cameras/)

Each JSON file specifies the geometry of one C-arm machine. All fields are required. The pipeline loads the config by machine name or file path.

JSON fields:

name (string) — human-readable machine name.

sid_mm (float) — Source-to-Image Distance in millimetres.

pixel_spacing_mm (float) — physical size of one detector pixel in mm.

image_size_px (int) — detector size in pixels (assumed square).

| File | Machine | SID (mm) | Pixel spacing (mm/px) | Source |
| --- | --- | --- | --- | --- |
| ziehm_vision_fd.json | Ziehm Vision RFD 3D | 1050 | 0.288 | DICOM tags (0018,1110) + (0019,1213) |
| siemens_cios_spin.json | Siemens Cios Spin | 1050 | 0.154 | Manufacturer specification |
| ge_oec_9900.json | GE OEC 9900 Elite | 1020 | 0.184 | Manufacturer specification |

IMPORTANT: Verify SID and pixel spacing against the real DICOM tags of your specific machine before running. The ziehm_vision_fd.json values were corrected from DICOM data (original config had SID=1110mm, pixel_spacing=0.2mm — both wrong).

# 7.  Setup and Installation

## 7.1  Prerequisites

| Component | Minimum | Recommended |
| --- | --- | --- |
| Python | 3.8 | 3.10 |
| GPU VRAM | 4 GB (CUDA) | 8 GB |
| RAM | 16 GB | 32 GB |
| Disk | 20 GB | 50 GB |
| OS | Windows 10 / Ubuntu 20.04 | Windows 11 / Ubuntu 22.04 |

## 7.2  Installation

# Clone the repository
git clone https://github.com/RaghuGangoluNITW/2D-3D-sci-registration
cd 2D-3D-sci-registration
git checkout release/v1.0

# Install CPU-only (testing only, very slow DRR rendering)
pip install -r requirements.txt

# Install with GPU support (CUDA 12.x) — STRONGLY RECOMMENDED
pip install torch==2.4.1 torchvision==0.19.1 --index-url https://download.pytorch.org/whl/cu121
pip install -r requirements.txt

## 7.3  Verify Installation

cd pipeline
python sanity_check_render.py
# Expected output: "DRR render test PASSED — saved sanity_drr.png"

## 7.4  Run DeepFluoro Test  (4.3 mm validated result)

Step 1: Download data

# Dataset: Grupp et al., Johns Hopkins 2020 (CC BY-NC 4.0)
# URL: https://doi.org/10.7281/T1/IFSXNV
# Place at: data/deepfluoro/ipcai_2020_full_res_data.h5

Step 2: Build DRR bank and run registration

cd pipeline
python build_drr_bank_deepfluoro.py    # ~5 minutes on GPU
python run_deepfluoro_multiview_test.py  # ~2 minutes
# Expected: "PDE = 4.3 mm  TARGET ACHIEVED"

## 7.5  Run VerSe Spine Test  (0.32 mm validated result)

# Download VerSe19 sub-verse010 from: https://osf.io/nqjyw/
# Place CT at: data/verse/sub-verse010_ct.nii.gz
cd pipeline
python run_verse_multiview_test.py
# Expected: "Mean PDE = 0.32 mm  TARGET ACHIEVED"

# 8.  Known Limitations in This Release

| Limitation | Impact | Planned Fix |
| --- | --- | --- |
| Spine landmark collinearity | AP vertebra centroids (L3-L5) near-collinear → PnP unstable for spine cases | Add radiopaque skin markers OR retrain CNN with pedicle detections |
| C-arm isocenter not independently calibrated | Camera position in CT space estimated, not measured | One-time phantom calibration session with C-arm engineer |
| CNN trained on few real X-rays (~30) | Landmark detection PCK@10mm = 44% AP / 8% LAT on real frames | Collect 13+ new patient datasets and retrain |
| No DICOM integration | Scripts read raw files, not directly from DICOM worklist | Planned for v1.1 |
| No real-time surgeon display | Registered pose is saved as image, not overlaid live | Planned for v1.2 |